MODIFY SLIDER #4 TO TILT TO 0 ORDER

(Note: Some pictures are of slider #5 which is mirror imaged from slider #4)

- Remove grating from slider #4.

- Remove cladding panel D8810-1. (Note: may also need to remove panels D8810-2 & D8801-2

- Remove panel D1346-1 from focal plane baffle assembly D1351-1. Note light seals near center of arc so they can be replaced when panel is reinstalled.

- Position slider #4 in front of grating box, not clamped.

- Tilt grating to reverse hard stop by hand by rotating grating tilt wheel.

- Check clearance between slider and grating box with mirror and feeler gauge. (Expect ~ 1/8 to ¼ in.)

- Remove tilt adjusting screw and tilt to stop.

- Check grating tilt with GUI.

- If tilt angle is -19.15 deg +- 0.05 deg OK, if greater replace with new one then adjust angle with adjusting screw, if less replace adjusting screw and adjust.

- Check clearance between slider and grating box with mirror and feeler gauge, if > ¼ in OK, continue.

- Replace limit switch brackets with new ones.

- Adjust secondary limit to 18.90 deg +- 0.1 deg.

- Adjust primary limit to 18.60+- 0.1 deg.

- Replace focal plane baffle panel, be sure light leaks are plugged.

- Replace cladding.