MODIFY SLIDER #4 TO TILT TO 0 ORDER

(Note: Some pictures are of slider #5 which is mirror imaged from slider #4)

- Remove grating from slider #4.

- Remove cladding panel D8810-1. (Note: may also need to remove panels D8810-2 & D8801-2

- Remove panel D1346-1 from focal plane baffle assembly D1351-1. Note light seals near center of arc so they can be replaced when panel is reinstalled.

- Position slider #4 in front of grating box, not clamped.

- Tilt grating to reverse hard stop by hand by rotating grating tilt wheel (D4581-1).

- Check clearance between slider and grating box with mirror and feeler gauge.

(Expect ~ 1/8 to ¼ in.)

7) Replace limit switch brackets with new ones (D4597-1, D4597-2).

- Remove tilt adjusting screw and tilt to stop (D4643-3).

- Check grating tilt with GUI.

- If tilt angle is -19.15 deg +- 0.05 deg OK, if greater replace with new one then adjust angle with adjusting screw.

  - if less replace adjusting screw and adjust

  - if greater replace with new one and adjust angle with adjusting screw. Check clearance of adjusting screw and brace D4370-1

- Check clearance between slider and grating box with mirror and feeler gauge, if > 1/8 in OK, continue.

- --- went to 7) ---- renumber ----

- Adjust secondary limit to 18.90 deg +- 0.1 deg.

- Adjust primary limit to 18.60+- 0.1 deg.

- Replace focal plane baffle panel, be sure light leaks are plugged.

- Replace cladding.