Deimos 6 month Preventative Maintenance

- Clean Slit Mask Cassette.  see to procedure

2) Drive Disk – inspect and clean friction surfaces and wipers

– check and adjust Patch and Skew  see procedure

3) Balance – check balance and record amperage readings. see procedure

4) Air Pressure – check all regulators and set pressure to value on label on gauge.

Grating System pressures are on Drawing D4102. Slitmask insertion             pressure is 40 psi and retraction pressure is 45 psi. Check system for leaks.

5) Grating system

Lead screws – clean and apply new lubricant with soft wipes.

Clean and grease chain.

Use - White Lithium Grease.

6) Inspect all optics. Look into camera from inside Deimos with imaging mirror in                         place and shutter open to inspect camera optics and see into dewar. Record quantity and position of any bubbles in coupling fluid.

7) Trolley system – check Deimos transport rollers, D6390 item 18, for patch.